## Slide 1
Product & System
User Guide
NEXT

## Slide 2
Basic Info
NEXT

## Slide 3
PANO, CEPH and CBCT imaging combined in a single device.
3 in 1 System
Low dose X-ray : This innovative technique allows patient to be protected by acquiring X-ray
images with low dose.
High-speed scan : With a 5.9 sec scan, this function reduces acquisition of low quality CT images caused by the movement of a patient during image capture.
UHD Mode : Acquire the PANO image with high resolution.
- Metal Artifact Reduction : Acquire the image by minimizing Metal Artifacts.
Outstanding Image Processing Algorithm
Composition & Features of the Product
PANO / CT Sensor Auto Switching System
CEPH Multi-FOV : Minimizes X-ray dose by selecting the size of the FOV upon the image capturing area.
Scan Ceph
OneShot Ceph
PANO
CT
NEXT

## Slide 4
Composition of Image Acquiring System
Console PC
Fiber Optic Cable(10m)
Warning system panel(Optional)
Warning Lamp
Door Interlock
Up/Down Switch
Ethernet Cable(OneShot CEPH)
RS-232
Ethernet Cable
Fiber Optic Cable
RS-232 Serial Cable(10m)
Exposure Switch
NEXT

## Slide 5
Structure, Function and Accessories of the Product
NEXT
CEPH device
Capture cephalometric image:
- Scan Type
- OneShot Type
Vertical Frame
Contains main control board and X-ray LED lamp.
Base
Base is for the support of the whole unit (Optional).
Column device
Adjust height of the unit.
Rotating Unit
Sensors and the X-ray Generator are embedded and rotates when capturing the PANO/CT image.
Handle Frame
Power switch and components needed for patient positioning.

## Slide 6
To reset the emergency stop switch, turn clockwise until it pops up.
Emergency Switch
The emergency switch is at the bottom part of the handle frame.
If there is any emergency situation that may harm the patient or damage the device, press the emergency switch immediately to stop the operation of the device.
LED turns on and the operation stops when pressing the emergency button.
The power switch is at the bottom part of the handle frame..
After turning on the switch, wait until the device initializes completely. Then execute the capture software.
Power Switch
Power Switch
Column Up/Down
Switch Connector
NEXT

## Slide 7
TOUCH LCD Button
① Touch Screen : Configure the parameter settings in each imaging mode.
② Laser Beam ON/OFF Button : Able to turn on/off laser beam.
③ READY Button : Click this button after completion of preparation for image capturing. (Parameter settings and patient positioning)
④ Chinrest/Frankfurt plane laser beam UP Button : For CT mode, you can move
the chinrest up. For PANO mode, you can move the Frankfurt plane laser beam up.
⑤ Column UP Button : Can move the Column up.
⑥ Chinrest/Frankfurt plane laser beam DOWN Button : For CT mode, you can move the chinrest down. For PANO mode, you can move the Frankfurt plane laser beam down.
⑦ Column DOWN Button : Can move the Column down.
1
2
3
4
5
6
7
By using the Column up/down button, control the height of the column and then proceed with patient positioning.
Column moves up by pressing the upper part of the button and
column moves down by pressing the bottom part of the button.
Column Up/Down
NEXT

## Slide 8
① Frankfurt Plane Laser Beam (Side) : PANO
② Mid-sagittal Plane Laser Beam (Front) : PANO/CT
③ Canine Beam (Side) : PANO
④ Upper CT Horizontal Laser Beam (Side) : For Large FOV
⑤ Lower CT Horizontal Laser Beam (Side) : For Small FOV
⑥ CT Vertical Laser Beam (Side)
Laser Beams
NEXT

## Slide 9
X-ray LED Lamp
When power supplies the device, the green LED light turns on and the LED light turns orange during X-ray exposure for image capturing.
Bite Block (Normal) & Temple Support
Place the head of the patient using the bite block and temple support for capturing the PANO image.
NEXT

## Slide 10
Nasal Positioner & Ear Rods
These accessories allow the patient’s head to be stable when capturing of the image in CEPH mode.
NEXT

## Slide 11
Exposure Switch
This switch is for image capturing.
The exposure switch allows the operator to control image acquisition from outside of the X-Ray room. Press and hold the exposure switch down until the acquisition is complete.
Exposure Button
Exposure
Indicator Light
Green : Ready
Orange: X-ray ON
If the exposure switch is pressed, the color of the exposure indicator light turns orange. It indicates that there is X-Ray exposure.
NEXT

## Slide 12
Bite block for acquiring
Sinus image.
Bite block for Edentulous patients.
Bite block for acquiring TMJ Open/Close image.
Chin Support : TMJ, Sinus, Edentulous
TMJ, Sinus, Edentulous.
Capture basic PANO image.
Chin Support w/ Bite Block (Normal)
Bite block for basic PANO image.
CARPUS Plate(CEPH)
Guide for capturing the carpus in CEPH mode.
NEXT

## Slide 13
Patient Registration
NEXT
EasyDent V4

## Slide 14
Execute the EasyDent V4 software by double clicking the EasyDent V4 icon.
EasyDent V4
NEXT
EasyDent4
Viewer
Click x 2!!

## Slide 15
Click the add patient icon. A window for registering a new patient will appear.
New Patient Registration
NEXT
Click!!
①
Fill in patient information in the blanks and then click the “Add” button to complete patient registration.
Click!!
②

## Slide 16
Gender
Select the gender of patient.
Date of Birth
Select the date of birth of patient.
Change the notation order from Settings.
Chart No.*
Enter the chart no. of the doctor.
The date registers automatically and you are able to cancel the set date from Settings. (Required)
SSN
Enter Social Security Number.
E-mail
Enter information of the patient in each section.
First Name*
Enter first name of patient.
You can change the notation order in Settings. (Required)
Last Name*
Enter last name of patient.
You can change the notation order in Settings. (Required)
Address
Enter address information.
Treatment
Select the state of treatment
Tel / Mobile
Enter telephone information.
NEXT

## Slide 17
Click!!
②
Click!!
①
Click!!
③
Click!!
④
Click!!
⑤
After updating patient info., click the “Modify” button to complete the modification.
Update Patient Info.
NEXT

## Slide 18
Patient Search(1/2)
NEXT
Click!!
①
Click!!
③
Click!!
④
Click!!
②
Click!!
⑤

## Slide 19
Patient Search(2/2)
NEXT
Double Click!!
①
Double-click on Chart No., First Name or Last Name to display virtual keyboard.

## Slide 20
How To Capture the Image
NEXT

## Slide 21
1
2
3
4
5
6
7
8
Touch LCD Parameter Setting (CBCT)
Function Description
① Display current information such as patient type and X-ray setting values.
② Select the FOV size of CT image.
③ Display the current selections in sequential order.
④ Select the CBCT / PANO / CEPH image capturing mode.
⑤ Enable the user to set a patient’s gender, X-ray intensity and kVp / mA.
⑥ Move to the previous step.
⑦ Move to the next step.
⑧ CONFIRM button.
NEXT

## Slide 22
Touch LCD Parameter Setting (CBCT)
1
2
① Select the position of the tooth.
② Display information on the selected tooth position.
Function Description
2
1
3
4
Function Description
① Tube voltage UP / DOWN button.
② Tube current UP / DOWN button.
③ Select the gender of the patient.
④ Able to select the X-ray intensity.
⑤ Displays the parameter settings.
NEXT

## Slide 23
Function Description
Touch Screen is automatically locked during exposure.
Function Description
2
1
① Widen the Temple Support.
② Narrow the Temple Support.
Function Description
Returns the Rotating device back to its initial position.
NEXT
Touch LCD Parameter Setting (CBCT)

## Slide 24
Touch LCD Parameter Setting (PANO)
Function Description
① Display Imaging parameter and Settings information.
② Select the Image Quality.
③ Select the dental arch of the patient.
④ Minimize image overlapping.
⑤ Display the current selections in sequential order.
1
2
3
5
4
NEXT

## Slide 25
Touch LCD Parameter Setting (CEPH)
Function Description
① Display information on the imaging parameter settings.
② Select the capturing type of the CEPH image.
③ Select the FOV size of the CEPH Image.
④ Adjust the X-ray exposure time.
⑤ Select the image quality.
1
2
3
4
Oneshot Ceph
Scan Ceph
5
NEXT

## Slide 26
Panorama Acquisition
Execute Capture Software (PANO)
Select the patient and then click the “Dental CT” button, and the capture software will execute.
NEXT
Click!!
①

## Slide 27
Image Capture Procedure - Standard Mode (1/7)
① Click the buttons in the following order. • PANO → HD(Optional) → Normal → Standard
Click!!
Click!!
If the patient age is under 14, child arch is selected automatically.
Minimizes teeth overlapping in the image.
Select when capturing in Bitewing mode.
The image is captured in the same order for the Right, Front, Left modes.
- AF function applied in PANO mode.
UHD : High resolution image - HD : High Definition Image
Normal : Normal Image
Panorama Acquisition
NEXT
② Click ‘CONFIRM’, then the device automatically turns to patient positioning mode.

## Slide 28
Image Capture Procedure - Standard Mode (2/7)
Panorama Acquisition
③ When the device stops moving, put the Chin Support with a normal Bite Block on the Chinrest.
④ For safety, the patient (especially child or regnant woman) should wear the lead apron.
⑤ Patient should take off all jewelry and metallic objects such as hair pins, spectacles, dentures and orthodontic
appliances. Place them on the accessory rack.
Lead Apron
Chin Support
w/ Bite Block
NEXT

## Slide 29
Image Capture Procedure - Standard Mode (3/7)
Panorama Acquisition
⑥ Bring patient to the device. Put a hygiene cover on the bite stick.
Hygiene Cover
⑧ Bring the patient to the device.
⑦ Adjust the temple supports, so that they fit snugly on both sides of the
patient’s head using the temple support wheel.
Temple Support
Button
Grab the handle tightly with two hands.
Chest leans against the device.
Move half a step forward.
Make sure the left and right shoulders remain level and the patient maintains this position.
Patient is to be positioned with the following procedures:
NEXT
⑨ Adjust the column height by pressing the column up/down switch (or button)
so that the patient’s chin reaches the chin rest.

## Slide 30
Image Capture Procedure - Standard Mode (4/7)
Panorama Acquisition
⑧ Check if the patient has bit the groove correctly with his/her upper and lower incisors.
⑨ Patient needs to be positioned correctly by checking the Mid-sagittal plane laser beam, Frankfurt plane laser beam
and the Canine laser beam .
Frankfurt Plane
Laser Beam
Canine Laser Beam
Adjusting Lever
Mid-sagittal Plane
Laser Beam
Mid-sagittal Plane
Laser Beam
Frankfurt Plane
Laser Beam
Canine Laser Beam
NEXT

## Slide 31
Image Capture Procedure - Standard Mode (5/7)
Panorama Acquisition
⑩ Position patient’s head with the temple support wheel so his/her head does not move.
⑪ Check the Mid-sagittal plane laser beam, Frankfurt plane laser beam and the Canine laser beam again.
Adjusting the width
of the Temple Support
Close eyes and remain still until the scanning is completed.
Keep tongue pressed against palate and breathe through nose.
Image capturing will begin.
Notify the following instructions to the patient.
NEXT

## Slide 32
Image Capture Procedure - Standard Mode (6/7)
Panorama Acquisition
⑪ Click the Ready button from the capture software.
⑫ Follow the instructions from the Guidance Message window, then proceed image capturing by pressing
the X-ray exposure switch.
Click!!
When pressing the X-ray exposure switch, the orange LED light turns on. This means that the device is exposing the X-ray.
NEXT

## Slide 33
Image Capture Procedure - Standard Mode (7/7)
⑬ After saving the acquired image, check whether it is saved in EasyDent V4 or not.
CASE #1. Normal / AF mode
Panorama Acquisition
NEXT
Click!!

## Slide 34
Image Capture Procedure - TMJ Mode (1/6)
① Click the buttons in the following order.
• PANO -> HD/Normal -> TMJ LAT Open
Click!!
② Click confirm and the device will move to the patient positioning mode.
Click!!
Capture the image with the same procedures as TMJ PA Open/Close mode.
* After completion of image acquisition in the TMJ LAT Open mode, the mode switches automatically to the TMJ LAT Close mode with the voice over instructions.
UHD : High resolution image - HD : High Definition Image
Normal : Normal Image
Panorama Acquisition
NEXT

## Slide 35
Image Capture Procedure - TMJ Mode (2/6)
Panorama Acquisition
③ Mount the TMJ(Sinus) bite block on the Chinrest and position the patient like the following image.
Chin Support:
TMJ, Sinus & Edentulous
Mid-sagittal Plane Laser Beam : Middle of the patient’s face
Frankfurt Plane Laser Beam : On the Frankfurt line
Canine Laser Beam : End of the lip line
Open
Close
NEXT

## Slide 36
Image Capture Procedure - TMJ Mode (3/6)
Panorama Acquisition
④ Fix the patient’s head with the temple support, so there is no movement.
⑤ Check the Mid-sagittal plane laser beam, Frankfurt plane laser beam and the Canine laser beam again.
Close eyes and remain still until the scanning is completed.
Keep your mouth wide open until the capturing is completed.
Notify the patient in advance that the capturing will be done twice (Open/Close)
Image capturing will begin.
Inform the patient of the following instructions:
⑥ Click READY from the capture software.
Click!!
NEXT

## Slide 37
Image Capture Procedure - TMJ Mode (4/6)
Panorama Acquisition
⑦ Proceed image capturing by pressing the exposure switch, as shown in the instructions in the Guidance Message Window.
Click!!
Click OK and the mode will be automatically switched to LAT Close mode.
Click!!
NEXT

## Slide 38
Image Capture Procedure - TMJ Mode (5/6)
Panorama Acquisition
⑩ Proceed image capturing by pressing the exposure switch as shown in the instructions in the Guidance Message Window.
⑧ Check the Mid-sagittal plane laser beam, Frankfurt plane laser beam and Canine laser beam again.
Close eyes and remain still until the scanning is completed.
Keep your mouth closed until the scanning is completed.
Image capturing will begin.
Notify these following instructions to the patient.
⑨ Click READY from the capture software.
Click!!
NEXT

## Slide 39
Image Capture Procedure - TMJ Mode (6/6)
⑪ After saving the acquired image, check whether it is saved in EasyDent V4 or not.
Panorama Acquisition
NEXT
Click!!
TMJ LAT Closed
TMJ LAT Open

## Slide 40
Image Capture Procedure - Sinus Mode (1/4)
① Click the buttons in the following order. • PANO -> HD/Normal -> Sinus LAT
Click!!
② Click CONFIRM, and then the device will turn to the patient positioning mode.
Click!!
In Sinus PA mode, capture the image with the same procedure.
UHD : High resolution image - HD : High Definition Image
Normal : Normal Image
Panorama Acquisition
NEXT

## Slide 41
Image Capture Procedure - Sinus Mode (2/4)
Panorama Acquisition
③ Mount the bite block on the Chinrest and position the patient like the following image.
Chin Support:
TMJ, Sinus & Edentulous
Mid-sagittal plane laser beam : on the center of the face of the patient
Frankfurt Laser Beam : on the tip of the nose when the patient’s head is bent back 10° ~ 15°
Canine Beam : corner of lip
NEXT

## Slide 42
Image Capture Procedure - Sinus Mode (3/4)
④ Position patient’s head with the temple support, so there is no movement.
⑤ Check the Mid-sagittal plane laser beam, Frankfurt plane laser beam and the Canine laser beam again.
Keep eyes closed and do not move until the scanning is completed.
Image capturing will begin.
Notify the following instructions to the patient:
⑥ Click READY from the capture software.
Click!!
Panorama Acquisition
NEXT

## Slide 43
Image Capture Procedure - Sinus Mode (4/4)
Panorama Acquisition
⑦ Proceed image capturing by pressing the exposure switch, as the instructions are displayed
on the capture software screen.
⑧ After saving the acquired image, check whether it is saved in EasyDent V4 or not.
NEXT
Click!!

## Slide 44
Cephalo Acquisition
NEXT
Execution of Capture Software (CEPH)
Select the patient, and then click “Dental CT”, capture software appears.
Click!!
①

## Slide 45
Image Capture Procedure - Lateral Mode (1/5)
① Click the buttons in the following order.
• CEPH → Lateral.
② Click CONFIRM, then the device will turn
to the patient positioning mode.
Click!!
Scan Time : 12.9 Sec
(Optional)
Scan Time : 13.5 Sec
In PA, SMV, Waters View and Carpus mode, capture the image with the same procedure.
Click!!
Cephalo Acquisition
NEXT

## Slide 46
Image Capture Procedure - Lateral Mode (2/5)
Cephalo Acquisition
③ If the device stops, widen the ear rods.
④ For safety, the patient (especially a child or pregnant woman) should wear a lead apron.
⑤ Patient should take off all jewelry and metallic objects such as hair pins, spectacles, dentures and orthodontic appliances and place them on the accessory rack.
Lead Apron
NEXT

## Slide 47
Image Capture Procedure - Lateral Mode (3/5)
Cephalo Acquisition
⑥ Bring the patient to the CEPH device.
⑦ Adjust the height of the device to suit the patient by pressing the column up/down button or switch (optional).
⑧ Position the ear rods to fit along the patient’s ear canals and adjust the nasal positioner.
⑨ Arrange the position of the nasal positioner, to the patient’s Nasion point.
NEXT

## Slide 48
Image Capture Procedure - Lateral Mode (4/5)
Cephalo Acquisition
⑨ Patient’s Frankfurt Plane should be parallel to the floor.
Do not move and keep your mouth and eyes closed until the image capturing is completed.
Image capturing will begin.
⑩ Notify the following instructions to the patient.
⑪ Click READY from the capture software.
Click!!
NEXT

## Slide 49
⑫ Proceed image capturing by pressing the exposure switch, as shown in the instructions displayed on the capture software screen.
Image Capture Procedure - Lateral Mode (5/5)
⑬ After saving the acquired image, check whether it is saved in EasyDent V4 or not.
Cephalo Acquisition
NEXT

## Slide 50
① Select the PA mode from the capture software and then click CONFIRM.
② Turn the ear rods 90° clockwise from their initial position.
④ Fold the nasal positioner.
③ Widen the distance between the two ear rods.
Click!!
⑤ Ask the patient to stand upright facing the sensor. Make sure that the patient’s shoulders are level and his/her neck is relaxed.
⑥ Click READY from the capture software and proceed to the image capture.
⑦ After saving the acquired image, check whether it is saved in EzDent-i or not.
NEXT
Cephalometric Image Capture Procedure
(In PA Mode)

## Slide 51
① Select the SMV mode from the capture software and click CONFIRM.
② Turn the ear rods 90° clockwise from their initial position.
④ Fold the nasal positioner.
③ Widen the distance between the two ear rods.
Click!!
⑤ Gently bend the patient’s head back until his/her Frankfurt plane is perpendicular to the floor as shown in image above.
⑥ Click READY from the capture software and proceed the image capture.
⑦ After saving the acquired image, check whether it is saved in EzDent-i or not.
NEXT
Cephalometric Image Capture Procedure
(In SMV Mode)

## Slide 52
① Select the Waters View mode from the capture software and click CONFIRM.
② Turn the ear rods 90° clockwise from their initial position.
④ Fold the nasal positioner
③ Widen the distance between the two ear rods.
⑤ Ask the patient to stand upright facing the sensor and bend his/her neck back about 35° - 40°.
⑥ Click READY from the capture software and proceed the image capture.
⑦ After saving the acquired image, check whether it is saved in EzDent-i or not.
Click!!
NEXT
Cephalometric Image Capture Procedure
(In Waters View Mode)

## Slide 53
① Select the Carpus mode from the capture software and click CONFIRM.
② Turn the ear rods 90° clockwise from their initial position.
④ Slide the locking bolt on the CARPUS plate towards the nasal positioner and turn it to the left.
③ Slide the CARPUS Plate onto the nasal positioner.
Click!!
⑤ Ask the patient to put his/her right hand flat on the CARPUS plate.
⑥ Click READY from the capture software and proceed the image capture.
⑦ After saving the acquired image, check whether it is saved in EzDent-i or not.
NEXT
Cephalometric Image Capture Procedure
(In CARPUS Mode)

## Slide 54
CT Acquisition
NEXT
Execution of Capture Software (CT)
Click!!
①
Select the patient, and then click “DENTAL CT” and the capture software displays.

## Slide 55
Image Capture Procedure - Standard Mode (1/7)
CT Acquisition
① Click the buttons in the following order.
• CT FOV > Position > Select the Image Quality > Metal Mode
② Click CONFIRM, then the device will turn to the patient positioning mode.
Click!!
Select the size of the CT FOV(Field of View).
- The FOV size is determined by the CT sensor attached on the device.
Select the ROI to be captured under VERTICAL POSITION:
- Mn(Mandible)
- Mx(Maxillary)
- Occl(Occlusal)
- TMJ(Temporomandibular Joint)
Select the ROI to be captured under HORIZONTAL POSITION.
- Right, Center, Left
5.9 second scan time.
Select either Standard or Application mode for voxel resolution.
Metal Artifact Reduction reduces appearance of scatter due to metal reflection.
NEXT

## Slide 56
Image Capture Procedure - Standard Mode (2/7)
CT Acquisition
③ When the device stops moving, put the Normal Bite Block on the Chinrest.
④ For safety, the patient (especially child or pregnant woman) should wear the lead apron
⑤ Patient should take off all jewelry and metallic objects such as hair pins, spectacles, dentures and orthodontic appliances
and place them on the accessory rack
Lead Apron
Bite Block : Normal
NEXT

## Slide 57
Image Capture Procedure - Standard Mode (3/7)
CT Acquisition
⑥ Put the hygiene cover on the bite stick
⑧ Bring the patient to the device.
⑦ Loosen the temple supports by turning the Temple Support Wheel.
Temple Support
Adjustment Button
Hygiene Cover
Grab the handle tightly with two hands.
Let the chest lean against the device.
Move half a step forward.
Left and right shoulder should remain level and keep holding the position.
Allow the patient to be positioned by following the procedures:
NEXT
⑨ Adjust the column height by pressing the column up/down switch (or button)
so that patient’s chin reaches the chin rest.

## Slide 58
Image Capture Procedure - Standard Mode (4/7)
CT Acquisition
⑩ Check if the patient has bit the groove correctly with his/her upper and lower incisors.
⑪ First, align the Mid-sagittal and CT vertical laser beam, and then align the CT horizontal laser beam by adjusting the height
of the chinrest.
Mid-sagittal Plane
laser beam
CT horizontal laser beam
CT vertical laser beam
Column DOWN
Button
Laser Beam On/Off
Column UP
button
NEXT

## Slide 59
Image Capture Procedure - Standard Mode (5/7)
CT Acquisition
⑫ Adjust the temple supports so they fit snugly on both sides of the patient’s head, using the temple support wheel.
⑬ Check the laser beam alignment again.
Adjust temple support
with Temple Support
Adjustment Button
Close eyes and remain still until the scanning is completed.
Image capturing will begin.
Notify the following instructions to the patient:
Mid-sagittal Plane
laser beam
CT horizontal laser beam
CT vertical laser beam
NEXT

## Slide 60
Image Capture Procedure - Standard Mode (6/7)
CT Acquisition
⑭ Click READY from the capture software.
⑮ Proceed image capturing by pressing the exposure switch as shown in the instructions, displayed on the capture software screen.
Click!!
When pressing the X-ray exposure switch, the orange LED light turns on. This means that the device is exposing the X-ray.
NEXT
Converting to DICOM file.

## Slide 61
Image Capture Procedure - Standard Mode (7/7)
CT Acquisition
NEXT
Click!!
①
Save the acquired image and check whether it is saved in EasyDent V4 or not.
Double click on CT name or thumbnail view to access the CT scan in the Ez3D software.
16

## Slide 62
Product Management
NEXT

## Slide 63
CASE #1. Device Not Operating
If any problem occurs while operating the device, perform the corresponding troubleshooting measures outlined in the table below. If the problem persists, please contact customer service.
CASE #2. X-ray Exposure Switch
Emergency Treatment
NEXT
CAUSE | SOLUTION
Fail to supply power | Check the power cable and the connection status.
Problem in initializing | Initialize the device completely and then try to operate again.
Fail in communication between PC & Unit | Check the connection status of serial port(RS232) which connects the PC to the device.
CAUSE | SOLUTION
Without pressing the “Ready” button | Check if the unit is completely ready for image capturing from the capture software.

## Slide 64
CASE #3. Problem in Image Capturing
CASE #4. Problem in Patient Positioning – Laser Beam Off
NEXT
CAUSE | SOLUTION
Fail in device initialization | Re-start to capture image after initialization of the unit. If this happens continuously, restart the unit and try again.
CAUSE | SOLUTION
Exceed the time limit for patient positioning | Turn the laser beam on again and proceed with patient positioning. Click the laser beam button from the unit or capture software. Then three laser beams will turn on.
Emergency Treatment

## Slide 65
Cleaning
  Please keep the device clean especially the parts where the patient is in contact with often such as the handle frame, chin support and the bite blocks.
  Do not use the spray cleaners or solvents from entering the device. It may cause damage to the electrical components.
  Do not use abrasive liquids such as acetone, gas, or oil. These may cause corrosion of the device’s surface.
  Do not use cleaning products which contain silicon.
  It may cause damage to the electrical components.
Before cleaning the device, make sure that the power of the device has been turned off completely and disconnect the power plug from the outlet.
NEXT
▣ ATTENTION

## Slide 66
▣ HOW TO CLEAN
NEXT
ACCESSORIES | HOW TO CLEAN
Bite Block | Before capturing the image of each patient, sterilize these accessories with alcohol and dry off with a soft cloth.
Temple Support | 
Chin Support (Normal / Sinus / TMJ) | 
All components that are in physical contact with the patient and the operator | 
PC and Peripheral Devices | Clean and keep them in good condition.
Outer part of the device | Wipe the outer part of the unit with a soft dry cloth everyday. Do not use any cleaners, bleaches or solvents for cleaning.
Cleaning

## Slide 67
Maintenance
Before inspection, ensure that the power of the device is off.
Do not open the cover of the device.
For the prevention of fire, please use the same fuse when replacing.
▣ ATTENTION
  Do not unplug the power plug by force.
  Do not keep the device and parts in a humid place.
  Keep the device and parts in an appropriate place to maintain them in good condition. They may be influenced by environmental factors such as temperature, lights, ventilation, dust, salt and so on.
  For items needed for image capturing, please arrange them and put them in proper places for the next image capture.
  Please check the ground connection of the device.
  Do not try to fix the device(including wires and cables) by yourself.
  It may cause accidents and damage to the device.
NEXT

## Slide 68
▣ Tasks for Maintenance
NEXT
Maintenance
CHECK LIST | Period
Before using the device, ensure that the device is clean and ready for use. Also, check the cleanliness of the parts where the patient is in contact with often. | Daily
After using the device, make sure that the main power of the switch has been turned off. | Daily
Ensure that the device is firmly plugged into a power supply device which is available only for the device. | Daily
Ensure that the power plug of the device is not unusually over heated. | Daily
While pressing the exposure switch(during X-ray exposure), please make sure that the indicator light turns on properly and remains orange to indicate there is no problem. | Daily
Check the proper operation of the X-ray exposure switch. | Daily
Check the status of the power cable. Ensure that there are no scratches or damage. | Weekly
Ensure that there is no problem on the operation of the emergency button. Entire operation of the device should stop if the emergency button is pressed. | Weekly
Ensure that all the labels attached on the device are clearly visible. | Monthly
Ensure that the voice over messages are audible during X-ray exposure. | Monthly

## Slide 69
Disposing Unit
How To Handle & Dispose Material
To reduce the environmental contamination, this dental x-ray device should be disposed by following the instructions. All parts and components which contain hazardous materials must be disposed in accordance with the disposal regulations.
Except the X-Ray tube, other parts are re-usable. (IEC 60601-1 6.8.2 j).
Please observe all regulations relevant to the disposal of waste in your country. For further information regarding the device disposal, please inquire at your customer service center.
This dental X-ray device must not be disposed of as regular
garbage.
Please clean, disinfect and sterilize the device before disassembling and disposing of its parts.
Classification | Main Material | Reusable | Special Treatment Needed (Mandatory) | Hazardous Waste (Needs to be handled separately)
Frame & Outer Cover | Aluminum & Plastic | ● |  | 
Motors |  | ● |  | 
Circuit boards |  | ● |  | 
Cable & Transformer | Copper | ● |  | 
 | Iron | ● |  | 
 | Oil |  | ● | 
Packing | Wood | ● |  | 
 | Cardboard | ● |  | 
 | Paper | ● |  | 
X-Ray Tube |  |  |  | ●
Sensor Head | Please return this part to the manufacturer, Vatech. |  |  | 
Other Parts |  |  | ● | 

## Slide 70
Thank you
FAQ