## Slide 1
Product & System
User Guide
NEXT

## Slide 2
Basic Info
NEXT

## Slide 3
PaX-i3D is an advanced digital dental diagnostic system that incorporates PANO, CEPH, and 3D CT imaging capabilities into a single system.
3 in 1 System
UHD Mode : Acquire the PANO image with high resolution.
AF Mode : Acquire a more accurate panoramic image by auto-focusing the capturing area.
Metal Artifact Reduction : Acquire the image by minimizing the Metal Artifacts.
Outstanding Image Processing Algorithm
Scan Ceph
OneShot Ceph
PANO
CT
PANO / CT Sensor Auto Switching System
CEPH Multi-FOV : Minimizes X-ray dose by selecting the size of the FOV upon the image capturing area.
NEXT
Composition & Features of the Product

## Slide 4
Composition of Image Acquiring System
Console PC
Fiber Optic Cable(10m)
Warning system panel(Optional)
Warning Lamp
Door Interlock
Exposure Switch
Ethernet Cable(OneShot CEPH)
RS-232
Ethernet Cable
Fiber Optic Cable
RS-232 Serial Cable(10m)
NEXT

## Slide 5
CEPH Unit
Capture cephalometric image:
- Scan Type
- OneShot Type
Vertical Frame
Contains main control board and X-ray LED lamp.
Base
Base is for the support of the whole unit (Optional).
Column Unit
Adjust height of the unit.
Rotating Unit
Sensors and the X-ray Generator are embedded and rotates when capturing the PANO/CT image.
Handle Frame
Power switch and the components needed for patient positioning.
Structure, Function and Accessories of the Product
NEXT

## Slide 6
To reset the emergency stop switch, turn it clockwise until it pops up.
Emergency Switch
The emergency switch is at the upper part of the handle frame.
If there is any emergency situation that may harm the patient or damage the device, press the emergency switch immediately to stop the operation of the device.
LED turns on and the operation stops when pressing the emergency button.
Power switch is at the bottom part of the handle frame..
After turning on the switch, wait until the device initializes completely. Then execute the capture software.
Power Switch
Power Switch
Column Up/Down
Switch Connector
NEXT

## Slide 7
HANDLE FRAME Button
① COLUMN : Adjust the height of the column when LED light is on.
② RETURN : Rotating unit returns to the initial position for the
next exposure.
③ / : Moves COLUMN/CHINREST up and down.
④ SELECT : COLUMN ↔ CHINREST mode button.
⑤ CHINREST : Adjust the height of the column when LED light is on.
1
2
3
4
5
By using the Column up/down button, control the height of the column and then proceed with patient positioning.
Column moves up by pressing the upper part of the button and
column moves down by pressing the bottom part of the button.
Column Up/Down
NEXT

## Slide 8
Bite Block (Normal) & Temple Support
Place the head of the patient using the bite block and temple support for capturing the PANO image.
X-ray LED Lamp
When power supplies the device, the green LED light turns on and the LED light turns orange during the X-ray exposure for image capturing.
NEXT

## Slide 9
Nasal Positioner & Ear Rods
These accessories allow the patient’s head to be stable when capturing the image in CEPH mode.
NEXT

## Slide 10
Exposure Switch
This switch is for image capturing.
The exposure switch allows the operator to control image acquisition from outside of the X-Ray room. Press and hold the exposure switch down until the acquisition is complete.
Exposure Button
Exposure
Indicator Light
Green : Ready
Orange: X-ray ON
If the exposure switch is pressed, the color of the exposure indicator light turns orange. It indicates that there is X-ray exposure.
NEXT

## Slide 11
Chin Support : Bite Block
Capture the basic PANO image.
Chin support: Sinus
Capture the sinus image.
Chin support : Edentulous
Chin support for edentulous patients.
Chin support: TMJ
Capture the TMJ Open/Close image.
Carpus Plate
Guide to capture the carpus in CEPH mode.
NEXT

## Slide 12
Patient Registration
NEXT
EasyDent V4

## Slide 13
Execute the EasyDent V4 software by double clicking the EasyDent V4 icon.
EasyDent V4
NEXT
EasyDent4
Viewer
Click x 2!!

## Slide 14
Click the add patient icon. A window for registering a new patient will appear.
New Patient Registration
NEXT
Click!!
①
Fill in patient information in the blanks and then click the “Add” button to complete patient registration.
Click!!
②

## Slide 15
Gender
Select the gender of patient.
Date of Birth
Select the date of birth of patient.
Change the notation order from Settings.
Chart No.*
Enter the chart no. of the doctor.
The date registers automatically and you are able to cancel the set date from Settings. (Required)
SSN
Enter Social Security Number.
E-mail
Enter information of the patient in each section.
First Name*
Enter first name of patient.
You can change the notation order in Settings. (Required)
Last Name*
Enter last name of patient.
You can change the notation order in Settings. (Required)
Address
Enter address information.
Treatment
Select the state of treatment
Tel / Mobile
Enter telephone information.
NEXT

## Slide 16
Click!!
②
Click!!
①
Click!!
③
Click!!
④
Click!!
⑤
After updating patient info., click the “Modify” button to complete the modification.
Update Patient Info.
NEXT

## Slide 17
Patient Search(1/2)
NEXT
Click!!
①
Click!!
③
Click!!
④
Click!!
②
Click!!
⑤

## Slide 18
Patient Search(2/2)
NEXT
Double Click!!
①
Double-click on Chart No., First Name or Last Name to display virtual keyboard.

## Slide 19
How To Capture the Image
NEXT

## Slide 20
NEXT
Execute Capture Software (PANO)
Select the patient and then click Dental CT,
then the capture software will execute.
Click!!
①

## Slide 21
Capture
Mode
Select the
capture mode of
PANO image.
Select
Modality
Select between
PANO and CEPH.
Close
Settings
Image Recon in Manual
Reconstruct the image manually by clicking this button if an error occurs during the automatic image reconstruction.
Phantom
Capture
Mode Display
Display the selected capture mode.
Patient
Info.
Display the chart no., name, gender, age and birthday of the patient.
Capture Status
& Image
Display the capture status in real time and view the image in advance.
DAP Value
Display the DAP
value.
Tube
Voltage/Current
Display the basic value of patients tube voltage and current. Finely adjust when needed.
Gender, Child and
Bone Density
Display the information on the registered patient automatically. Edit when necessary.
Guidance Message
Display the various directions and guidance through voice over messages.
X-ray Exposure
The color of the sign turns yellow when there is X-ray exposure.
NEXT

## Slide 22
Image Capture Procedure - Standard Mode 1/7
① Click the buttons in the following order. • PANO → HD(Optional) → Normal → Standard.
Click!!
② Click ‘CONFIRM’, then the device automatically turns to patient positioning mode.
Click!!
If the patient is under the age of 14, child arch is selected automatically.
This mode minimizes teeth overlapping in the image.
This arch is selected when capturing in the Bitewing mode.
The image is captured in the same order for the Right, Front, Left modes.
- AF function applied in PANO mode.
UHD : High resolution image - HD : High Definition Image
Normal : Normal Image
NEXT

## Slide 23
Image Capture Procedure - Standard Mode 2/7
③ When the device stops moving, put the Chin Support w/ Bite Block on the Chinrest.
④ For safety, the patient (especially child or pregnant woman) should wear the lead apron.
⑤ Patient should take off all jewelry and metallic objects such as hair pins, spectacles, dentures and orthodontic appliances
and place them on the accessory rack.
Lead Apron
Chin Support
w/ Bite Block
NEXT

## Slide 24
Image Capture Procedure - Standard Mode 3/7
⑥ Bring the patient to the device. Put a hygiene cover on the bite stick.
Hygiene Cover
⑧ Bring the patient to the device.
Grab the handle tightly with two hands.
Chest leans against the device
Move half a step forward..
Make sure the left and right shoulders remain level and the patient maintains this position.
Patient is to be positioned with the following procedures:
⑦ Adjust the temple supports so that they fit snugly on both sides of the patient’s head using the temple support wheel.
Temple Support
Wheel
⑨ Adjust the column height by pressing the column up/down switch (or button)
so that the patient ‘s chin reaches the chin rest.
NEXT

## Slide 25
Image Capture Procedure - Standard Mode 4/7
⑨ Patient needs to be positioned correctly by checking the Mid-sagittal plane laser beam,
Frankfurt plane laser beam and the Canine laser beam .
Frankfurt Plane
Laser Beam
Adjusting Lever
Canine Laser Beam
Adjusting Lever
Mid-sagittal Plane
Laser Beam
Mid-sagittal Plane
Lase Beam
Frankfurt Plane
Laser Beam
Canine Laser Beam
⑧ Check if the patient has bit the groove correctly with his/her upper and
lower incisors.
NEXT

## Slide 26
Image Capture Procedure - Standard Mode 5/7
⑩ Position patient’s head with the temple support wheel so his/her head does not move..
⑪ Check the Mid-sagittal plane laser beam, Frankfurt plane laser beam and Canine laser beam again.
Close eyes and remain still until the scanning is completed.
Keep tongue pressed against the palate and breathe through the nose.
Image capturing will begin.
Notify the following instructions to the patient.
Temple Support
Wheel
NEXT

## Slide 27
⑪ Click the Ready button from the capture software.
⑫ Follow the instructions from the Guidance Message window, then proceed image capturing by pressing the X-ray exposure switch.
Click!!
Image Capture Procedure - Standard Mode 6/7
When pressing the X-ray exposure switch, the orange LED light turns on. This means that the device is exposing the X-ray.
NEXT

## Slide 28
Image Capture Procedure - Bitewing Mode
① Click the buttons in the following order. • PANO -> HD-> Orthogonal -> Bitewing
② Click confirm and the device will turn to the patient positioning mode.
Click!!
◎ The next step is the same for the image capturing procedure in standard mode.
In Bitewing Right, Bitewing Incisor and Bitewing Left modes, same procedure is done to capture images.
If the patient is under the age of 14, child arch is selected automatically.
This mode minimizes teeth overlapping in the image.
Select this orthogonal arch when capturing the image in Bitewing mode.
UHD : High resolution image - HD : High Definition Image
Normal : Normal Image
Click!!
NEXT

## Slide 29
Image Capture Procedure - TMJ Mode 1/6
① Click the buttons in the following order.
• PANO -> HD/Normal -> TMJ LAT Open
Click!!
② Click confirm and the device will turn to the patient positioning mode.
Click!!
Capture the image with the same procedure as in TMJ PA Open/Close mode.
* After completion of image acquisition in TMJ LAT Open mode, the mode switches automatically to TMJ LAT Close mode with the voice over instructions.
UHD : High resolution image - HD : High Definition Image
Normal : Normal Image
NEXT

## Slide 30
Image Capture Procedure - TMJ Mode 2/6
③ Mount the TMJ bite on the Chinrest and position the patient like the following image.
TMJ Bite
NEXT
Mid-sagittal plane laser beam : vertical line through center of patient’s face
Frankfurt Laser Beam : superior tragus to infra-orbital rim
Canine Beam : corner of lip

## Slide 31
Image Capture Procedure - TMJ Mode 3/6
④ Fix the patient’s head with the temple support, so there is no movement.
⑤ Check the Mid-sagittal plane laser beam, Frankfurt plane laser beam and the Canine laser beam again.
Close eyes and remain still until the scanning is completed.
Keep your mouth wide open until the capturing is completed.
Notify to the patient in advance that the capturing will be done twice (Open/Close)
Image capturing will begin.
Inform the patient of the following instructions:
⑥ Click READY from the capture software.
Click!!
Temple Support Wheel
NEXT

## Slide 32
⑦ Proceed image capturing by pressing the exposure switch shown in the instructions in the Guidance Message Window.
Image Capture Procedure - TMJ Mode 4/6
Click!!
Click OK and the mode will be automatically switched to LAT Close mode.
Click!!
NEXT

## Slide 33
Image Capture Procedure - TMJ Mode 5/6
⑧ Check the Mid-sagittal plane laser beam, Frankfurt plane laser beam and the Canine laser
beam again.
⑨ Click READY from the capture software.
Click!!
⑩ Proceed image capturing by pressing the exposure switch as shown in the instructions in the Guidance Message Window.
Close eyes and remain still until the scanning is completed.
Keep your mouth closed until the scanning is completed.
Image capturing will begin.
Notify these following instructions to the patient.
NEXT

## Slide 34
⑬ After saving the acquired image, check whether it is saved in EasyDent V4 or not.
Image Capture Procedure - TMJ Mode 6/6
NEXT
TMJ LAT Closed
TMJ LAT Open
Click!!

## Slide 35
Image Capture Procedure - Sinus Mode 1/4
① Click the buttons in the following order. • PANO -> HD/Normal -> Sinus LAT
Click!!
② Click CONFIRM, then the device will turn to the patient positioning mode.
Click!!
In Sinus PA mode, capture the image with the same procedure.
UHD : High resolution image - HD : High Definition Image
Normal : Normal Image
NEXT

## Slide 36
Image Capture Procedure - Sinus Mode 2/4
③ Mount the Sinus bite on the Chinrest and position the patient like the image below
Sinus Bite
Mid-sagittal plane laser beam : vertical line through center of patient’s face
Frankfurt Laser Beam : line through from tip of nose to top of ear
Canine Beam : corner of lip
NEXT

## Slide 37
Image Capture Procedure - Sinus Mode 3/4
④ Position patient’s head with the temple support, so there is no movement.
⑤ Check the Mid-sagittal plane laser beam, Frankfurt plane laser beam and the Canine laser beam again.
Keep your eyes closed and do not move until the scanning is completed.
Image capturing will begin.
Notify the following instructions to the patient.
⑥ Click READY from the capture software.
Click!!
Temple Support Wheel
NEXT

## Slide 38
⑦ Proceed image capturing by pressing the exposure switch, as the instructions are displayed
on the capture software screen.
Image Capture Procedure - Sinus Mode 4/4
⑧ After saving the acquired image, refresh patient to view captured image.
Click!!
NEXT

## Slide 39
Select the patient, and then click Dental Ct, capture software appears.
NEXT
Execution of Capture Software (CEPH)
Click!!
①

## Slide 40
Image Capture Procedure - Lateral Mode 1/6
① Click the buttons in the following order.
• CEPH → Lateral.
② Click CONFIRM, then the device will turn
to the patient positioning mode.
Click!!
Click!!
Scan Time : 12.9 Sec
In PA, SMV, Waters View and Carpus mode, capture the image with the same procedure.
(Optional)
Scan Time : 13.5 Sec
NEXT

## Slide 41
Image Capture Procedure - Lateral Mode 2/6
③ If the device stops, widen the ear rods.
④ For safety, patient (especially child or pregnant woman) should wear the lead apron.
⑤ Patient should take off all jewelry and metallic objects such as hair pins, spectacles, dentures and
orthodontic appliances and place them on the accessory rack.
Lead Apron
Accessory Rack
NEXT

## Slide 42
Image Capture Procedure - Lateral Mode 3/6
⑥ Bring the patient to the CEPH unit.
⑦ Adjust the height of the unit to suit the patient by pressing the column up/down button or switch (optional).
NEXT

## Slide 43
Image Capture Procedure - Lateral Mode 4/6
⑧ Position the ear rods to fit along the patient’s ear canals and adjust the nasal positioner
⑨ Position the nasal positioner on the patient’s Nasion point.
NEXT

## Slide 44
Image Capture Procedure - Lateral Mode 5/6
⑨ Patient’s Frankfurt Plane should be parallel to the floor.
⑩ Notify the following instructions to the patient.
Do not move and keep your mouth and eyes closed until the image capturing is completed.
Image capturing will begin.
⑪ Click READY from the capture software.
NEXT
Click!!

## Slide 45
Image Capture Procedure - Lateral Mode 6/6
⑫ Proceed image capturing by pressing the exposure switch, as shown in the instructions in the Guidance Message Window.
⑬ After saving the acquired image, refresh patient to view captured image.
NEXT
Click!!

## Slide 46
① Select the PA mode from the capture software and then click CONFIRM.
② Turn the ear rods 90° clockwise from their initial position.
④ Fold the nasal positioner.
③ Widen the distance between the two ear rods.
Click!!
⑤ Ask the patient to stand upright facing the sensor. Make sure that the patient’s shoulders are level and his/her neck is relaxed.
⑥ Click READY from the capture software and proceed to the image capture.
⑦ After saving the acquired image, check whether it is saved in EzDent-i or not.
NEXT
Cephalometric Image Capture Procedure
(In PA Mode)

## Slide 47
① Select the SMV mode from the capture software and click CONFIRM.
② Turn the ear rods 90° clockwise from their initial position.
④ Fold the nasal positioner.
③ Widen the distance between the two ear rods.
Click!!
⑤ Gently bend the patient’s head back until his/her Frankfurt plane is perpendicular to the floor as shown in image above.
⑥ Click READY from the capture software and proceed the image capture.
⑦ After saving the acquired image, check whether it is saved in EzDent-i or not.
NEXT
Cephalometric Image Capture Procedure
(In SMV Mode)

## Slide 48
① Select the Waters View mode from the capture software and click CONFIRM.
② Turn the ear rods 90° clockwise from their initial position.
④ Fold the nasal positioner
③ Widen the distance between the two ear rods.
⑤ Ask the patient to stand upright facing the sensor and bend his/her neck back about 35° - 40°.
⑥ Click READY from the capture software and proceed the image capture.
⑦ After saving the acquired image, check whether it is saved in EzDent-i or not.
Click!!
NEXT
Cephalometric Image Capture Procedure
(In Waters View Mode)

## Slide 49
① Select the Carpus mode from the capture software and click CONFIRM.
② Turn the ear rods 90° clockwise from their initial position.
④ Slide the locking bolt on the CARPUS plate towards the nasal positioner and turn it to the left.
③ Slide the CARPUS Plate onto the nasal positioner.
Click!!
⑤ Ask the patient to put his/her right hand flat on the CARPUS plate.
⑥ Click READY from the capture software and proceed the image capture.
⑦ After saving the acquired image, check whether it is saved in EzDent-i or not.
NEXT
Cephalometric Image Capture Procedure
(In CARPUS Mode)

## Slide 50
Execution of Capture Software (CT)
Select the patient, and then click the “DENTAL CT” button and capture software displays.
NEXT
Click!!
①

## Slide 51
Image Capture Procedure - Standard Mode (1/7)
① Click the buttons in the following order.
• CT FOV > Position > Select the Image Quality > Metal Mode
② Click CONFIRM, then the device will turn to the patient positioning mode.
Click!!
Select the size of the CT FOV(Field of View).
- The FOV size is determined by the CT sensor attached on the device.
Select the ROI to be captured under the VERTICAL POSITION.
- Mn(Mandible)
- Mx(Maxillary)
- Occl(Occlusal)
- TMJ(Temporomandibular Joint)
Select the ROI to be captured under the HORIZONTAL POSITION.
- Right, Center, Left
Select the image acquisition speed under High (24 s) or Standard (15 s).
Select either Standard or Application mode for voxel resolution.
Metal Artifact Reduction reduces scatter caused by metal reflection.
NEXT

## Slide 52
Image Capture Procedure - Standard Mode (2/7)
③ When the device stops moving, put the Chin Support w/ Bite Block on the Chinrest.
④ For safety, the patient (especially child or pregnant woman) should wear the lead apron.
⑤ Patient should take off all jewelry and metallic objects such as hair pins, spectacles, dentures and orthodontic appliances
and place them on the accessory rack.
Lead Apron
Chin Support
w/ Bite Block
NEXT

## Slide 53
⑥ Bring the patient to the device. Put the hygiene cover on the bite stick.
Hygiene Cover
⑧ Bring the patient to the device.
⑦ Loosen the temple supports by turning the Temple Support Wheel .
Temple Support
Wheel
Image Capture Procedure - Standard Mode (3/7)
Grab the handle tightly with two hands.
Let the chest lean against the device.
Move half a step forward..
Left and right shoulders remain level and keep holding the position.
Allow the patient to be positioned by following the procedures:
⑨ Adjust the column height by pressing the column up/down switch (or button)
so that the patient’s chin reaches the chin rest.
NEXT

## Slide 54
⑧ Check if the patient has bit the groove correctly with his/her upper and lower
incisors.
⑨ First, align the Mid-sagittal, CT vertical laser beam and then align the CT horizontal laser beam by adjusting the height of the chinrest.
Image Capture Procedure - Standard Mode (4/7)
Mid-sagittal Plane
laser beam
CT horizontal laser beam
CT vertical laser beam
CHINREST LED Light
Chin rest
Up/Down button
NEXT

## Slide 55
⑩ Adjust the temple supports so they fit snugly on both sides of the patient’s head, using the temple support wheel.
⑪ Check the beam alignment again.
Temple
Support
Wheel
Image Capture Procedure - Standard Mode (5/7)
Close the eyes and remain still until the scanning is completed.
Image capturing will begin.
Notify the following instructions to the patient.
NEXT

## Slide 56
⑫ Click READY from the capture software.
⑬ Proceed image capturing by pressing the exposure switch as shown in the instructions in the Guidance Message Window.
Click!!
When pressing the X-ray exposure switch, the orange LED light turns on. This means that the device is exposing the X-ray.
Image Capture Procedure - Standard Mode (6/7)
Capturing Image
image reconstruction is in progress
Converting to DICOM file.
NEXT

## Slide 57
⑭ After saving the acquired image, refresh patient to view captured image.
Double click to view the CT scan in the Ez3D plus software.
Image Capture Procedure - Standard Mode (7/7)
NEXT
Click!!
①

## Slide 58
Product Management
NEXT

## Slide 59
CASE #1. Device Not Operating
If any problem occurs while operating the device, perform the corresponding troubleshooting measures outlined in the table below. If the problem persists, please contact customer service.
CASE #2. X-ray Exposure Switch
Emergency Treatment
NEXT
CAUSE | SOLUTION
Fail to supply power | Check the power cable and the connection status.
Problem in initializing | Initialize the device completely and then try to operate again.
Fail in communication between PC & Unit | Check the connection status of serial port(RS232) which connects the PC to the device.
CAUSE | SOLUTION
Without pressing the “Ready” button | Check if the unit is completely ready for image capturing from the capture software.

## Slide 60
CASE #3. Problem in Image Capturing
CASE #4. Problem in Patient Positioning – Laser Beam Off
NEXT
CAUSE | SOLUTION
Fail in device initialization | Re-start to capture image after initialization of the unit. If this happens continuously, restart the unit and try again.
CAUSE | SOLUTION
Exceed the time limit for patient positioning | Turn the laser beam on again and proceed with patient positioning. Click the laser beam button from the unit or capture software. Then three laser beams will turn on.
Emergency Treatment

## Slide 61
Cleaning
  Please keep the device clean especially the parts where the patient is in contact with often such as the handle frame, chin support and the bite blocks.
  Do not use the spray cleaners or solvents from entering the device. It may cause damage to the electrical components.
  Do not use the abrasive liquids such as acetone, gas, or oil. These may cause corrosion of the device’s surface.
  Do not use cleaning products which contain silicon.
  It may cause damage to the electrical components.
Before cleaning the device, make sure that the power of the device has been turned off completely and disconnect the power plug from the outlet.
NEXT
▣ ATTENTION

## Slide 62
▣ HOW TO CLEAN
NEXT
ACCESSORIES | HOW TO CLEAN
Bite Block | Before capturing the image of each patient, sterilize these accessories with alcohol and dry off with a soft cloth.
Temple Support | 
Chin Support (Normal / Sinus / TMJ) | 
All components that are in physical contact with the patient and the operator | 
PC and Peripheral Devices | Clean and keep them in good condition.
Outer part of the device | Wipe the outer part of the unit with a soft dry cloth everyday. Do not use any cleaners, bleaches or solvents for cleaning.
Cleaning

## Slide 63
Maintenance
Before inspection, ensure that the power of the device is off.
Do not open the cover of the device.
For the prevention of fire, please use the same fuse when replacing.
▣ ATTENTION
  Do not unplug the power plug by force.
  Do not keep the device and parts in a humid place.
  Keep the device and parts in an appropriate place to maintain them in good condition. They may be influenced by environmental factors such as temperature, lights, ventilation, dust, salt and so on.
  For items needed for image capturing, please arrange them and put them in proper places for the next image capturing.
  Please check the ground connection of the device.
  Do not try to fix the device(including wires and cables) by yourself.
  It may cause accidents and damage to the device.
NEXT

## Slide 64
▣ Tasks for Maintenance
NEXT
Maintenance
CHECK LIST | Period
Before using the device, ensure that the device is clean and ready for use. Also, check the cleanliness of the parts where the patient is in contact most often. | Daily
After using the device, make sure that the main power of the switch has been turned off. | Daily
Ensure that the device is firmly plugged into a power supply device which is available only for the device. | Daily
Ensure that the power plug of the device is not unusually over heated. | Daily
While pressing the exposure switch(during X-ray exposure), please make sure that the indicator light turns on properly and remains orange to indicate there is no problem. | Daily
Check the proper operation of the X-ray exposure switch. | Daily
Check the status of the power cable. Ensure that there are no scratches or damage. | Weekly
Ensure that there is no problem on the operation of the emergency button. Entire operation of the device should be stopped if the emergency button is pressed. | Weekly
Ensure that all the labels attached on the device are clearly visible. | Monthly
Ensure that the voice over messages are audible during X-ray exposure. | Monthly

## Slide 65
Disposing Unit
How To Handle & Dispose Material
To reduce the environmental contamination, this dental x-ray device should be disposed by following the instructions. All parts and components which contain hazardous materials must be disposed in accordance with the disposal regulations.
Except the X-Ray tube, other parts are re-usable. (IEC 60601-1 6.8.2 j).
Please observe all regulations relevant to the disposal of waste in your country. For further information regarding the device disposal, please inquire to your customer service center.
This dental X-ray device must not be disposed of as regular
garbage.
Please clean, disinfect and sterilize the device before dissembling and disposing of its parts.
Classification | Main Material | Reusable | Special Treatment Needed (Mandatory) | Hazardous Waste (Needs to be handled separately)
Frame & Outer Cover | Aluminum & Plastic | ● |  | 
Motors |  | ● |  | 
Circuit boards |  | ● |  | 
Cable & Transformer | Copper | ● |  | 
 | Iron | ● |  | 
 | Oil |  | ● | 
Packing | Wood | ● |  | 
 | Cardboard | ● |  | 
 | Paper | ● |  | 
X-Ray Tube |  |  |  | ●
Sensor Head | Please return this part to the manufacturer, Vatech. |  |  | 
Other Parts |  |  | ● | 

## Slide 66
Thank you
FAQ