## Slide 1
Product & System
User Guide
NEXT

## Slide 2
Click
One Section
Basic
Info
Image Capturing
Procedure
Product
Management
Patient
Registration

## Slide 3
COPYRIGHT© 2014 VATECH Global.
ALL RIGHTS RESERVED
Basic Info
NEXT

## Slide 4
Digital PANO(basic) and CEPH(option) functions combined in one device
2 in 1 System
PANO UDH Mode : Acquire the PANO Image with high resolution.
AF Mode (Optional) : Acquire the accurate PANO image by using the Auto-focusing mode in the capturing area.
PANO Image Algorithm of the Next Generation
Composition & Features of the Product
NEXT

## Slide 5
Composition of Image Acquisition System
NEXT
Up/Down Switch
Exposure Switch
Warning System Panel
Warning Lamp
Door Interlock
RS232 Serial Cable 10M/32.8’
Ethernet LAN Cable 10M/32.8’
Frame Grabber(LVDS) Cable 7M/23’
OneShot Ceph
Ethernet LAN Cable 10M/32.8’

## Slide 6
CEPH Unit
Unit for acquiring the CEPH image (Optional).
Vertical Frame
Contains main control board and the X-ray LED lamp.
Base
Base is for the support of the whole unit (Optional).
Column Unit
Height of the device is adjusted.
Rotating Unit
Sensor and X-ray Generator rotates when capturing the PANO image.
Handle Frame
Power switch and the components needed for patient positioning.
Structure, Function and Accessories of the Product
NEXT

## Slide 7
Turn on the switch and wait until the product initializes completely. Then, execute the Capture Software.
Power Switch
To re-operate the device, please rotate the button clockwise, in the
direction of the arrow.
Emergency Switch
If any emergency situation occurs, please press the emergency switch to stop the operation of the device.
NEXT

## Slide 8
Up/Down Button
By using the Column up/down button, control the height of the column and then proceed with patient positioning.
Column moves up by pressing the button.
Column moves down by pressing the button.
By using the Column up/down button, control the height of the column and then proceed with patient positioning.
Column moves up when pressing the upper part of the switch and column goes down when pressing the lower part of the switch.
Up/Down Switch
NEXT

## Slide 9
Bite Block (Normal) & Temple Support
Place the head of the patient using the bite block and temple support for capturing the PANO image.
X-ray LED Lamp
If power supplies the device, the green LED light turns on and the LED light changes to orange during the X-ray exposure for image capturing.
COPYRIGHT© 2014 VATECH Global.
ALL RIGHTS RESERVED
NEXT

## Slide 10
Nasal Positioner & Ear Rods
These accessories help keep the patient’s head stable and to not move when capturing the image in CEPH mode.
NEXT

## Slide 11
Exposure Switch
This switch is for image capturing.
The exposure switch allows the operator to control image acquisition from outside of the X-Ray room. Press and hold the exposure switch down until the acquisition is complete.
Exposure Button
Exposure
Indicator Light
Green : Ready
Orange : X-ray ON
If the exposure switch is pressed, the color of the exposure indicator light turns orange.
This indicates that there is X-Ray exposure.
NEXT

## Slide 12
Chin Support w/ Bite Block
Capturing the basic PANO image.
Chin support: Sinus
Capturing the sinus image.
Chin support : Edentulous
Chin support for edentulous patients.
Chin support: TMJ
Capturing the TMJ Open/Close image.
Carpus Plate
Guide for capturing the carpus in CEPH mode.
NEXT

## Slide 13
Patient Registration
NEXT
EasyDent V4

## Slide 14
Execute the EasyDent V4 software by double clicking the EasyDent V4 icon.
EasyDent V4
NEXT
EasyDent4
Viewer
Click x 2!!

## Slide 15
Click the add patient icon. A window for registering a new patient will appear.
New Patient Registration
NEXT
Click!!
①
Fill in patient information in the blanks and then click the “Add” button to complete patient registration.
Click!!
②

## Slide 16
Gender
Select the gender of patient.
Date of Birth
Select the date of birth of patient.
Change the notation order from Settings.
Chart No.*
Enter the chart no. of the doctor.
The date registers automatically and you are able to cancel the set date from Settings. (Required)
SSN
Enter Social Security Number.
E-mail
Enter information of the patient in each section.
First Name*
Enter first name of patient.
You can change the notation order in Settings. (Required)
Last Name*
Enter last name of patient.
You can change the notation order in Settings. (Required)
Address
Enter address information.
Treatment
Select the state of treatment
Tel / Mobile
Enter telephone information.
NEXT

## Slide 17
Click!!
②
Click!!
①
Click!!
③
Click!!
④
Click!!
⑤
After updating patient info., click the “Modify” button to complete the modification.
Update Patient Info.
NEXT

## Slide 18
Patient Search(1/2)
NEXT
Click!!
①
Click!!
③
Click!!
④
Click!!
②
Click!!
⑤

## Slide 19
Patient Search(2/2)
NEXT
Double Click!!
①
Double-click on Chart No., First Name or Last Name to display virtual keyboard.

## Slide 20
How to Capture the Image
NEXT

## Slide 21
First, select a patient, click the “Panorama” button and the capture software will execute.
Execute Capture Software (PANO)
NEXT
Click!!
①

## Slide 22
Capture Mode
Select image quality, capturing area and special mode.
Settings & Option Button
Choose from buttons for phantom mode, image recon and settings..
Mode Display
Displays the selected capture mode.
Indicate X-ray Exposure
The color of the sign turns yellow during X-ray exposure.
Patient Info.
Display the chart no., name, gender, age and birthday of the patient.
Guide Message
Displays various directions and guidance with messages.
X-ray Exposure Condition
Select the gender or child, as well as the bone density depending on patient condition.
NEXT

## Slide 23
Panoramic Image Capture Procedure
(Standard Mode)
Click buttons in the following order.
• PANO → HD → Normal → Standard
Click ‘CONFIRM’ button. Then the device will
move to the patient positioning mode.
NEXT
Click!!
①
Click!!
②

## Slide 24
Lead Apron
For safety, help patient wear the Lead Apron.
Tell patient to take off all jewelry and metallic objects such as hair pins, spectacles, dentures and orthodontic appliances. Place them on the accessory rack.
NEXT
Click!!
①
Click!!
②
Panoramic Image Capture Procedure
(Patient Positioning)

## Slide 25
Chin Support
w/ Bite Block
Mount the “Chin Support w/ Bite Block” on the Chinrest.
Use this chin
support for
edentulous patients.
Chin support :
Edentulous
Hygiene Cover
Cover the bite stick with the Hygiene cover.
Hygiene covers are consumables. You are able to purchase these hygiene covers from your regional CS center.
Hygiene Cover
NEXT
Panoramic Image Capture Procedure
(Patient Positioning)

## Slide 26
Panoramic Image Capture Procedure
(Patient Positioning)
Patient moves to the inside of the device.
Adjust the height of the column by pressing the up/down switch (or button) until the patient‘s chin reaches the chin rest.
Grab the handle of the device tightly with two hands.
Move half step forward and let chest lean against the device.
Make sure left and right shoulders remain level and keep this position.
Follow patient procedures for proper image capture.
NEXT

## Slide 27
Panoramic Image Capture Procedure
(Adjusting the Laser Beam)
Check if the patient bit on the groove correctly, with his/her upper and lower incisors.
Arrange the position of the patient’s head by checking the Mid-sagittal plane laser beam
Arrange the position of the patient’s head so it is level by checking the Frankfurt Plane laser beam.
Frankfurt Plane Laser Beam
Adjusting
Lever
By adjusting the Canine laser beam, position the patient’s canine properly.
Canine Laser Beam
Adjusting Lever
NEXT
Click!!
①
Click!!
②
Click!!
③

## Slide 28
Close eyes and remain still until the scanning is completed.
Keep your tongue pressed against the palate and breathe through the nose.
Image capture will begin.
Notify the following instructions to the patient.
Temple Support
Wheel
Keep patient’s head stable by adjusting the temple support.
Check the Mid-sagittal plane laser beam, Frankfurt plane laser beam and the Canine laser beam again.
Panoramic Image Capture Procedure
(Patient Positioning - Head)
NEXT

## Slide 29
Panoramic Image Capture Procedure
(In Standard Mode)
Click the “READY” button and follow the instructions indicated on the capture software. Then complete the image capture by pressing the x-ray exposure switch.
READY
OK
NEXT

## Slide 30
After saving the acquired image, check whether it is saved in EasyDent V4 or not.
CASE #1. Basic option.
CASE #2. In case AF or UHD option is applied.
NEXT
Panoramic Image Capture Procedure
(In Standard Mode)
Click!!
Click!!

## Slide 31
Panoramic Image Capture Procedure
(In Bitewing Mode)
Click the buttons in the following order.
• PANO → HD→ Orthogonal → Bitewing
CONFIRM
READY
Capture the image with the same procedure as the Bitewing Right, Bitewing Incisor, and Bitewing Left modes.
The capture procedure is the same as the standard mode in the next step.
NEXT

## Slide 32
Panoramic Image Capture Procedure
(In TMJ Open Mode)
CONFIRM
READY
Chin support: TMJ
OK
After completion of image acquisition in the TMJ LAT Open mode, the mode switches automatically to TMJ LAT Close mode with a voice over message.
Capture the image with the same procedure in
TMJ PA Open/Close mode.
Click the buttons in the following order.
• PANO → HD→ TMJ LAT Open
NEXT

## Slide 33
Panoramic Image Capture Procedure
(In TMJ Close Mode)
Chin support: TMJ
READY
CONFIRM
SAVE
TMJ LAT Open
TMJ LAT Close
Click CONFIRM.
NEXT

## Slide 34
Panoramic Image Capture Procedure
(In Sinus Mode)
Capture procedure is the same as the standard mode in the next step .
CONFIRM
Chin support: Sinus
READY
Capture the image with the same procedure
in Sinus PA mode.
Click the buttons in the following order.
• PANO → HD→ Sinus LAT
NEXT

## Slide 35
After selecting a patient, click the “Cephalo” button, and the capture software will execute.
NEXT
Execute Capture Software
(CEPH)
Click!!
①

## Slide 36
Click the buttons in the following order.
• CEPH → Lateral
Cephalometric Image Capture Procedure
(In Lateral Mode)
(Optional)
Click CONFIRM, then the device will
turn to the patient positioning mode.
Click!!
②
Click!!
①
NEXT

## Slide 37
Place the ear rods to fit along the patient’s ear canals, so the patient’s head will not move during capture of the image.
Place the nasal positioner on the patient’s Nasion point.
NEXT
Cephalometric Image Capture Procedure
(Patient Positioning)

## Slide 38
Click the “READY” button and follow the instructions indicated on the capture software. Then complete the image capture by pressing the x-ray exposure switch.
READY
OK
NEXT
Cephalometric Image Capture Procedure
(In Lateral Mode)

## Slide 39
① Select the PA mode from the capture software and then click CONFIRM.
② Turn the ear rods 90° clockwise from their initial position.
④ Fold the nasal positioner.
③ Widen the distance between the two ear rods.
Click!!
⑤ Ask the patient to stand upright facing the sensor. Make sure that the patient’s shoulders are level and his/her neck is relaxed.
⑥ Click READY from the capture software and proceed to the image capture.
⑦ After saving the acquired image, check whether it is saved in EzDent-i or not.
NEXT
Cephalometric Image Capture Procedure
(In PA Mode)

## Slide 40
① Select the SMV mode from the capture software and click CONFIRM.
② Turn the ear rods 90° clockwise from their initial position.
④ Fold the nasal positioner.
③ Widen the distance between the two ear rods.
Click!!
⑤ Gently bend the patient’s head back until his/her Frankfurt plane is perpendicular to the floor as shown in image above.
⑥ Click READY from the capture software and proceed the image capture.
⑦ After saving the acquired image, check whether it is saved in EzDent-i or not.
NEXT
Cephalometric Image Capture Procedure
(In SMV Mode)

## Slide 41
① Select the Waters View mode from the capture software and click CONFIRM.
② Turn the ear rods 90° clockwise from their initial position.
④ Fold the nasal positioner
③ Widen the distance between the two ear rods.
⑤ Ask the patient to stand upright facing the sensor and bend his/her neck back about 35° - 40°.
⑥ Click READY from the capture software and proceed the image capture.
⑦ After saving the acquired image, check whether it is saved in EzDent-i or not.
Click!!
NEXT
Cephalometric Image Capture Procedure
(In Waters View Mode)

## Slide 42
① Select the Carpus mode from the capture software and click CONFIRM.
② Turn the ear rods 90° clockwise from their initial position.
④ Slide the locking bolt on the CARPUS plate towards the nasal positioner and turn it to the left.
③ Slide the CARPUS Plate onto the nasal positioner.
Click!!
⑤ Ask the patient to put his/her right hand flat on the CARPUS plate.
⑥ Click READY from the capture software and proceed the image capture.
⑦ After saving the acquired image, check whether it is saved in EzDent-i or not.
NEXT
Cephalometric Image Capture Procedure
(In CARPUS Mode)

## Slide 43
Product Management
NEXT

## Slide 44
CASE #1. Device Not Operating
If any problem occurs while operating the device, perform the corresponding troubleshooting measures outlined in the table below. If the problem persists, please contact customer service.
CASE #2. X-ray Exposure Switch
Emergency Treatment
NEXT
CAUSE | SOLUTION
Fail to supply power | Check the power cable and the connection status.
Problem initializing | Initialize the device completely and then try to operate again.
Fail in communication between PC & Unit | Check the connection status of serial port(RS232) which connects the PC to the device.
CAUSE | SOLUTION
Without pressing the “Ready” button | Check if the unit is completely ready for image capture from the capture software.

## Slide 45
CASE #3. Problem in Image Capturing
CASE #4. Problem in Patient Positioning – Laser Beam Off Issue
NEXT
CAUSE | SOLUTION
Fail in device initialization | Re-start to capture image after initialization of the unit. If this happens continuously, restart the unit and try again.
CAUSE | SOLUTION
Exceed the time limit for patient positioning | Turn the laser beam on again and proceed to patient positioning. Click the laser beam button from the unit or capture software, then three laser beams will turn on.
Emergency Treatment

## Slide 46
Cleaning
  Please keep the device clean, especially the part where the patient touches or has contact with often such as the handle frame, chin support and the bite blocks.
  Do not use the spray cleaners or solvents to clean the device. It may cause damage on the electrical components.
  Do not use abrasive liquids such as acetone, gas, or oil. It may cause corrosion of the device’s surface.
  Do not use cleaning products which contain silicon.
  It may cause damage on the electrical components.
Before cleaning the device, make sure that the power of the device has been turned off completely and disconnect the power plug from the outlet.
NEXT
▣ ATTENTION

## Slide 47
▣ HOW TO CLEAN
NEXT
ACCESSORIES | HOW TO CLEAN
Bite Block | Before capturing the image of each patient, sterilize these accessories with alcohol and dry off with a soft cloth.
Temple Support | 
Chin Support (Normal / Sinus / TMJ) | 
All components that are in physical contact with the patient and the operator | 
PC and Peripheral Devices | Clean and keep them in good condition.
Outer part of the Device | Wipe the outer part of the unit with a soft dry cloth everyday. Do not use any cleaners, bleaches or solvents for cleaning.
Cleaning

## Slide 48
Maintenance
Before inspection, ensure that the power of the device is turned off.
Do not open the cover of the device.
For the prevention of fire, please use the same fuse when replacing.
▣ ATTENTION
  Do not unplug the power plug with force.
  Do not keep the device and parts in a humid place.
  Keep the device and parts in an appropriate place to maintain them in a good condition. They may be influenced by environmental factors such as temperature, lights, ventilation, dust, salt and etc.
  For items needed for image capturing, please arrange them and put them in proper places for the next image capture.
  Please check the ground connection of the device.
  Do not try to fix the device(including wires and cables) by yourself.
  This may cause accidents and damage to the device.
NEXT

## Slide 49
▣ Tasks for Maintenance
NEXT
Maintenance
CHECK LIST | Period
Before using the device, ensure that the device is clean and ready for use. Also, check the cleanliness of the areas where the patient is in contact with often. | Daily
After using the device, make sure that the main power of the switch has been turned off. | Daily
Ensure that the device is firmly plugged into a power supply device, which is available only for the device. | Daily
Ensure that the power plug of the device is not unusually over heated. | Daily
While pressing the exposure switch (during X-ray exposure), please make sure that the indicator light turns on properly and remains orange, indicating there is no problem. | Daily
Check the proper operation of the x-ray exposure switch. | Daily
Check the status of the power cable. Ensure that there are no scratches or damage. | Weekly
Ensure that there is no problem with the operation of the emergency button. Entire operation of the device should stop if the emergency button is pressed. | Weekly
Ensure that all the labels attached on the device are clearly visible. | Monthly
Ensure that the voice over messages are audible during X-Ray exposure. | Monthly

## Slide 50
Disposing Unit
How To Handle & Dispose Material
To reduce environmental contamination, this dental x-ray device should be disposed of by following these instructions. All parts and components which contain hazardous materials must be disposed of in accordance with the disposal regulations.
Except the X-Ray tube, all other parts are re-usable. (IEC 60601-1 6.8.2 j).
Please observe all regulations relevant to the disposal of waste in your country. For further information regarding device disposal, please inquire at your customer service center.
This dental X-ray device must not be disposed of as regular
garbage.
Please clean, disinfect and sterilize the device before disassembling and disposing of its parts.
Classification | Main Material | Reusable | Special Treatment Needed (Mandatory) | Hazardous Waste (Needs to be handled separately)
Frame & Outer Cover | Aluminum & Plastic | ● |  | 
Motors |  | ● |  | 
Circuit boards |  | ● |  | 
Cable & Transformer | Copper | ● |  | 
 | Iron | ● |  | 
 | Oil |  | ● | 
Packing | Wood | ● |  | 
 | Cardboard | ● |  | 
 | Paper | ● |  | 
X-Ray Tube |  |  |  | ●
Sensor Head | Please return these parts to the manufacturer, Vatech. |  |  | 
Other Parts |  |  | ● | 

## Slide 51
Thank you
FAQ